User Manual-QL-800-O3

Thanks for purchasing our product. Please kindly read the user manual carefully before using.

Keys Operation:

(Note: The long press function only work in detection mode. When you change or

reset any data, please remember to click“ON/OFF” button to confirm and save the settings. Remember to turn on the pump when start the detection. Please turn off the pump when there is no need to detect, which will be helpful for the service life of the pump)

Power On

Press the “ON/OFF” for three seconds, the detector will be turned on and the pump will start to work. The screen displays the following interface successively: Brand and Logo, main parameter interface, initialing countdown. The countdown would be 60 or 120 seconds, which is to make sure that the sensors are fully activated. After countdown, the device enter normal detection interface.

Zero Calibration

Zero calibration must be done in clean air. Better to do once after alarming happens.

Press ”MENU” to enter parameter setting mode, select the "calibrate zero", and press “ON/OFF” to enter the mode. Observe the value of “real time concentration”. If the reading is not zero and the drifting is too big when the reading is stable, then it is necessary to conduct zero point calibration.

Press “ON/OFF”, the real-time concentration will go back to zero, and press ”ESC” to go back to detection mode, gas concentration values also back to zero.

Note: Zero calibration of oxygen detector, nitrogen detector and carbon dioxide detector (target gases are component of the air) can not be done in the air. Only the target point calibration (value of target point concentration is the standard concentration values in the clean air) can be done.

Process of Over-range incorrect operation

Users should avoid using gases which exceed the measure range to impact the sensor, because it will affect the lifespan and sensitivity of the sensor, even, "poison" the sensor. If there is any over-range incorrect operation makes the detector displaying an concentration at a large reading, the remedy is to take the device out of the environment immediately and put it in clean air for over half an hour, and then observe the reading, if it keeps going down, then wait until the reading back to zero before powering it off, and do the zero point calibration next time before using it. If the reading maintained at full scale, user should send the device back to manufacturer or agent for repair or replacing the sensor.

① — Air Inlet

② — Air Outlet

③ — air chamber of sensor

④ — LCD

⑤ — PLUS

⑥ — MINUS

⑦—MENU

⑧— ESC

⑨— ON/OFF

Target point calibration

Note: Calibration needs to be done by professionals and with necessary accessories and standard gas.

Prepare the standard gas before starting the target point calibration. In detection mode, press ”MENU”, select “calibrate span”, press ON/OFF button to enter. Then you will need to enter password. (If you need password, please contact the manufacturer to get it)

After entering the password successfully, press ON/OFF to enter the target point calibration interface, as shown in the picture. The interface will display “concentration”, “set value. The concentration is the result got from the calculation of several internal calibration parameters. The process of calibration is the process to recalculate the internal parameters of the device. The set value is the value need to be entered manually; it should be the concentration of the standard gas to be calibrated. Select the parameter under the current interface, modify its value with “PLUS” and “MINUS” button.

The calibration process is as below:

- Connect the detector with standard gas cylinder by hose.

- Enter the target point calibration interface, set the “SET value” as the standard gas concentration.

- Release the standard gas to the detector in a flow rate of 500mL/min. The value of “concentration” will become larger gradually, and become stable after 30seconds.

- When the value of “concentration” is stable, press the “ON/OFF” button, the value of “concentration” will be same with or almost same with the “Set value”. The calibration is finished.

- Cut off the standard gas supply. Press “ESC” to go back to detection interface.